## Slide 1: Aerospace Symposium 2021
Composites Radius Inspection – Calibration and defect sizing
Tom Bertenshaw – Senior NDT Research Engineer
Thursday 22nd April 2021
Uncontrolled data - The documents within this transmission contain NO technical data subject to the "UK Export Control Act 2002 or EU Dual-Use Regulation (Council Regulation (EC) No 428/2009), as amended"

## Slide 2
Aerospace Symposium 2021
  Composites Radius Inspection – Calibration and defect sizing
    The presentation covers:
      The applicable specifications
      The parameters evaluated
      A brief description of a few radius inspection techniques
      Proposed blocks to evaluate the parameters
      Methods to size defects
Export Classification - NSR

## Slide 3
Aerospace Symposium 2021
  Composites Radius Inspection – Calibration and defect sizing
    Inspection Composite components with phased array UT is commonplace in the industry.
    This has led to best practises for equipment testing, characterisation and annual calibration. Specifications include:
      BS EN ISO 18563-1:2015 - Non-destructive testing – Characterization and verification of ultrasonic phased array equipment. Part 1: Instruments
      BS EN ISO 18563-2:2017 - Non-destructive testing – Characterization and verification of ultrasonic phased array equipment. Part 2: Probes
      BS EN ISO 18563-3:2015 - Non-destructive testing – Characterization and verification of ultrasonic phased array equipment. Part 3: Combined Equipment
      ASTM E317: Evaluating performance characteristics of Ultrasonic Pulse-echo Testing Instruments and Systems with the use of electronic measurement systems
Export Classification - NSR

## Slide 4
Aerospace Symposium 2021
  Composites Radius Inspection – Calibration and defect sizing
    Parameters subject to evaluation (not necessarily aligned with specs shown on previous page)
Category | Parameter
Probe | Spectrum & Bandwidth
 | Pulse shape
 | Sensitivity
Instruments | Pulser
 | Frequency response
 | Accuracy of calibrated attenuator/gain control
 | Vertical linearity
Combined equipment | Resolution
 | Sensitivity and noise
 | Backlash
 | Sizing of defects
 | Overlapping
 | Normalisation
Subject to deviation from current requirements due to radius geometry
    Types of Composite radius inspection
Radius Tan-line to tan-line coverage using curved array probe and water box
Radius tan line to tan-line coverage using flat array probe utilising solid wedge and beam steering
Radius Tan-line to tan-line coverage using curved array probe
partial radius coverage
using curved array probe
partial radius coverage
using flat array probe
Export Classification - NSR

## Slide 5
Aerospace Symposium 2021
  Composites Radius Inspection – Calibration of parameters
    The spectrum & bandwidth, pulse shape and sensitivity are evaluated using the same “half pipe” cal block
Note: the radius of the calibration block must match that of the radius probe
Probe
Rexolite ®
calibration
block
Export Classification - NSR

## Slide 6
Aerospace Symposium 2021
  Composites Radius Inspection – Calibration of parameters
    A number of blocks are proposed here.
    For vertical linearity
    To calibrate resolution and sensitivity, the whole system must be considered. Therefore the probe is evaluated with the instrument (and shoe where relevant)
Note: the radius of the calibration block (R1) must match that of the radius probe
¾ R1
R1
Rexolite ®
calibration
block
Probe
Far surface resolution block
Internal radii:
Far surface resolution block
Entry surface resolution block
Entry surface resolution block
External radii:
Internal radius sensitivity block
External radius sensitivity block
Export Classification - NSR

## Slide 7
Aerospace Symposium 2021
  Composites Radius Inspection – Calibration of parameters
    Backlash and overlapping
    Normalisation
    Defect Sizing
Overlap region subject to backlash
Angled faces of the saw tooth profile
Straight edge perpendicular to direction of inspection
45°
A
Section A-A
Internal radius surface
A
External radius surface
5mm typ.
For defect measurement in radii, there are two methods proposed:
sizing by comparison to reference standard
sizing by geometrical calculation
Export Classification - NSR

## Slide 8
Aerospace Symposium 2021
  Composites Radius Inspection – Calibration of parameters: example
    Internal radius inspection at a GKN site
    Rexolite® block designed in this case for Normalisation, Sensitivity and Resolution
Sensitivity block
Interface surface resolution block
Normalisation and far surface resolution block
Export Classification - NSR

## Slide 9
Aerospace Symposium 2021
  Composites Radius Inspection – Defect sizing by comparison to reference standard
    This requires a reference standard to have a number of defects placed through the thickness of the laminate
    This example can be used for internal and external radii
    Inserted defects only – no flat bottomed hole defects
    Sizing the defect is a scaling task. This can be done by the following steps:
      measured on the part and the depth of the defect recorded
      The defect on the reference standard or reference sample is found with a corresponding depth (or as close a possible) and the artificial defect is measured
      The defect (yactual) can then be calculated
6x6mm PTFE (or similar) insert defects are spaced through the laminate
Aperture direction
Where:
ypart = the measured defect on the part in the aperture direction
yref stnd = the measured defect on the reference standard in the aperture direction
Export Classification - NSR

## Slide 10
Aerospace Symposium 2021
  Composites Radius Inspection – Defect sizing by geometrical calculation
    For internal radii, a defect is calculated by two factors:
      Ratio of arc lengths of the probe and the part
      Beam spread due to the changing radius size as the ultrasound progresses through the part
rP
r1
r2
rP – Radius of Probe
r1 – Internal radius of part
α
r1
r2
d2
d
6mm defect (in red)
Where:
d - defect size (6mm in this case, or measured defect)
r1 – Radius of part or distance to front surface / interface
r2 - defect depth (Rdef)
d2 – scaled defect (from beam spread)
R1
rP – Radius of Probe
r1 – Internal radius of part
AP – Arc length of probe active area
A1 – Arc length of internal radius
R1
AP
A1
AP
A1
Export Classification - NSR

## Slide 11
Aerospace Symposium 2021
  Composites Radius Inspection – Applying calculations
Arc length = 56.5mm
Arc length = 45.5mm
C-scan
Active aperture direction
    This example shows the measurement of the defects in the active aperture direction before & after in a reference standard.
    Although there is an improvement, ±2mm is not really good enough.
Arc length = 33.6mm
Area covered from IML surface
Export Classification - NSR

## Slide 12
Aerospace Symposium 2021
  Composites Radius Inspection – Summary
    A number of test blocks are proposed for the calibration of a number phased array radius inspection techniques
    Proposals for the calculation of defect size for internal radius inspection has been made
    Ongoing work: Calculation of defect size for external radius inspection
Export Classification - NSR